Azide Blood Agar Base

Azide Blood Agar Base is used for the selective isolation and cultivation of Staphylococcus and Streptococcus species from mixed bacterial flora.

| REF: V/AZ01.100 100 Gram REF: V/ AZ01.500 500 Gram | REF: V/AZ01.250 250 Gram |
| --- | --- |

## CLINICAL SIGNIFICANCE

## Azide Blood Agar Base is recommended for the isolation and cultivation of Streptococcus species from clinical and nonclinical specimens. It is a modification of the broth medium originally formulated by Edwards for the detection of Streptococci from bovine mastitis cases (1). The original broth medium of Edwards was modified to a blood agar by Packer containing sodium azide and crystal violet (2).

## METHOD PRINCIPLE

## Peptone special and HM Peptone B are the sources of carbon, nitrogen and essential growth factors. Sodium azide acts as a selective agent by suppressing the growth of gram-negative bacteria. It also prevents the swarming of Proteus (3, 4). Sodium chloride helps to maintain the osmotic balance of the medium. The media can be supplemented with sterile defibrinated blood to prepare blood agar. Blood serves as an additional source of growth factors and it also helps to visualize the haemolytic pattern. The pH of the medium influences the inhibitory action of sodium azide (2). At pH 7.2, sodium azide does not interfere with the haemolytic reactions of Streptococci ; however, haemolytic pattern of Streptococci is different on Azide Blood Agar as compared on nonselective blood agar. For best results, use light inoculum and incubate anaerobically for enhancement in haemolytic reaction. Different types of haemolysis can be visualized on blood agar plates (5).

## MEDIA COMPOSITION

| ## Item | ## Formula per liter of medium |
| --- | --- |
| ## Peptone, special ## Beef extract ## Sodium chloride ## Sodium azide ## Agar | ## 10.00 gm. ## 3.000 gm. ## 5.000 gm. ## 0.200 gm. ## 15.00 gm. |

## Final pH 7.2 ± 0.2 at 25°C

## PRECAUTIONS AND WARNINGS

## Media to be handled by entitled and professionally educated person. Do not ingest or inhale.

## Good Laboratories practices using appropriate precautions should be followed in:

## Wearing personnel protective equipment (overall, gloves, glasses,).

## Do not pipette by mouth.

## In case of contact with eyes or skin; rinse immediately with plenty of soap and water. In case of severe injuries; seek medical advice immediately.

## Respect country requirement for waste disposal.

## S56: dispose of this material and its container at hazardous or special waste collection point.

## S57: use appropriate container to avoid environmental contamination.

## S61: avoid release in environment.

## For further information, refer to the Azide Blood Agar Base Material safety data sheet.

## EQUIPMENT REQUIRED NOT PROVIDED

## Sterile plates

## Incubator

## Autoclave

## STORAGE AND STABILITY

Lab.Vie Azide Blood Agar base should be stored between 10-30°C in a firmly closed container and the prepared medium at 2-8°C. Use before expiry date on the label. On opening, product should be properly stored dry, after tightly capping the bottle in order to avoid lump development due to the hygroscopic nature of the product. Improper storage of the product may lead to lump formation. Store in a dry ventilated area protected from extremes of temperature and sources of ignition. Seal the container tightly after use. Product performance is best if used within stated expiry period.

## PREPARATION

## Suspend 33.2 grams in 1000 ml of distilled water.

## Heat to boiling to dissolve the medium completely.

## Sterilize by autoclaving at 15 lbs. pressure (121°C) for 15 minutes.

## Cool to 45-50°C. For preparing Blood Agar plates, 5% v/v sterile defibrinated blood is added aseptically.

## Mix well and pour into sterile Petri plates.

## Deterioration

## The color of Lab.Vie Azide Blood Agar Base is Cream to yellow homogeneous free flowing powder. Prepared Basal Medium is Yellow coloured, clear to slightly opalescent gel. After addition of 5%w/v sterile defibrinated blood is Cherry red coloured, opaque gel forms in Petri plates, which darkens on standing. If there are any physical changes for powder or signs of deterioration (shrinking, cracking, or discoloration), and contaminations for hydrated media, discard the medium.

## SPECIMEN

## Clinical samples – Blood

## QUALITY CONTROL

## To ensure adequate quality control, it is recommended that positive and negative control included in each run. If control values are found outside the defined range, check the system performance. If control still out of range please contact Lab.Vie technical support.

## PERFORMANCE CHARACTERISTICS

| Organism | Growth | Haemolysis |
| --- | --- | --- |
| Enterococcus faecalis ATCC 29212 | luxuriant | alpha/gamma |
| Escherichia coli ATCC 25922 | None -poor | none |
| Staphylococcus epidermidis ATCC 12228 | luxuriant | none |
| Streptococcus pyogenes ATCC 19615 | Inhibited | Beta |
| Streptococcus pneumoniae ATCC 6303 | Good | Alpha |

## REFERENCES

## Edwards, 1933, J. Comp. Pathol. Therap., 46:211.

## Packer, 1943, J. Bacteriol., 1943, 46:343

## Snyder and Lichstein, 1940, J. Infect. Dis., 67:113.

## Lichstein and Snyder, 1941, J. Bacteriol., 42:653. 5.

## Isenberg, (Ed.), 1992, Clinical Microbiology Procedures Handbook, Vol. I., American Society for Microbiology, Washington, D.C.

## Murray P. R., Baron E. J., Jorgensen J. H., Pfaller M. A., Yolken R. H., (Eds.),8th Ed., 2003, Manual of Clinical Microbiology, ASM, Washington, D.C

| ## SYMBOLS IN PRODUCT LABELLING |  |  |
| --- | --- | --- |
| IVD | For in-vitro diagnostic use | Number of <n> test in the pack |
| LOT | Batch Code/Lot number | Caution |
| REF | Catalogue Number | Do not use if package is damaged |
|  | Temperature Limitation | Consult Instruction for use |
|  | Expiration Date |  |
|  | Manufactured by |  |

Ismailia – Free zone, Ismailia – Egypt                                                                                    IFU-S-02, Rev. 03 - December 2019

Post code-41511

E-mail   : admin@labvielab.com

Website: www.labvielab.com